## Calculation
Procedure
1 | Default starting point
| Default R-curve as set in the GUI |  |  | SpO2 (DUT) used = | -28.149 | R^2 + | -10.293 | R + | 107.36
2 | Run the test procedure - enter measured SP02 values in the yellow field
| Obtain the new DUT R-curve |  |  |  | Preset | Measured |  | Calculated
[5 empty] | SpO2 WhaleTeq | SpO2 DUT |  | R DUT
[5 empty] | 100 | 98 |  | 0.4221 |  | Measured SP02 values
[5 empty] | 95 | 94 |  | 0.52994
[5 empty] | 90 | 87 |  | 0.68707
[5 empty] | 85 | 83 |  | 0.76523
[5 empty] | 80 | 78 |  | 0.85469
[5 empty] | 75 | 74 |  | 0.92105
[5 empty] | 68 | 65 |  | 1.0574
3 | Use new R-Curve values before running the validation test
| Type in the reference and measured SpO2 values |  |  | SpO2 (DUT) new = | -25.3259 | R^2 + | -13.4324 | R + | 110.0096 | New R-Curve Values
[5 empty] | Preset | Measured |  | deviation | Pass/Fail | Fail
[5 empty] | SpO2 WhaleTeq | SpO2 DUT |  | in % | < 3.5% | Pass
[5 empty] | 100 | 99.8274 |  | 0.1726% | Pass | Corrected Calculated SP02 values
[5 empty] | 95 | 95.7787 |  | -0.8197% | Pass
[5 empty] | 90 | 88.8253 |  | 1.305% | Pass
[5 empty] | 85 | 84.9004 |  | 0.1172% | Pass
[5 empty] | 80 | 80.0286 |  | -0.03579% | Pass
[5 empty] | 75 | 76.153 |  | -1.537% | Pass
[5 empty] | 68 | 67.4866 |  | 0.755% | Pass
4 | Run the validation procedure - enter measured SP02 values in the yellow field
Obtain the new DUT R curve
[5 empty] | Set of Measured | Measured |  | deviation | Pass/Fail
[5 empty] | SpO2 reference | SpO2 DUT |  | in % | < 3.5%
[5 empty] | 100 | 97 |  | 3% | Pass | Measured SP02 values
[5 empty] | 95 | 95 |  | 0% | Pass
[5 empty] | 90 | 92 |  | -2.222% | Pass
[5 empty] | 85 | 83 |  | 2.353% | Pass
[5 empty] | 80 | 78 |  | 2.5% | Pass
[5 empty] | 75 | 74 |  | 1.333% | Pass
[5 empty] | 68 | 66 |  | 2.941% | Pass
THEN HUMAN SUBJECT TESTING NEW TAB.
Explanation of calculations
|  | 1 | Use the DUT measured SpO2 values
|  | 2 | Use the R curve of the DUT
|  | 3 | Calculate the R values the DUT determined
|  | 4 | Use these R values and the SpO2 va;lues form the reference
|  | 5 | Use polynominal trendline as new R curve

## Report
6024 Silver Creek Valley Road
San Jose, CA 95138
Date | Jan. 27, 2021
Test report
Within the limitations of the equipement and procedure used to obtain the SpO2 reference
and DUT SpO2 values below (OB1203SD-RL-EVK evaluation board, WhaleTeq AECG100
tester, PPG-2R 940 test head, WhaleTeq to DUT light cone), the results indicate that the
DUT most liekly will pass human subject testing.
| Set of Measured | Measured | deviation
| SpO2 reference | SpO2 DUT | in %
Pass/Fail
< 3.5%
| 100 | 97 | 3%
| 95 | 95 | 0%
| 90 | 92 | -2.222%
| 85 | 83 | 2.353%
| 80 | 78 | 2.5%
| 75 | 74 | 1.333%
| 68 | 66 | 2.941%
_________________________
Name